## Slide 1: Preparing Your Point of Care Testing Program for a CAP Inspection
Adrienne M. Malta, MBA, MT(ASCP), Sr. Manager, Inspection Services
College of American Pathologists
March 20, 2014
cap.org
©2014 College of American Pathologists. All rights reserved.

## Slide 2: Learning Objectives:
Understand how to best prepare your POCT for an accreditation inspection
Identify the most common deficiencies that are cited and how to avoid them
Update on new requirements for 2014

## Slide 3: What IS point-of-care-testing?
Non-dedicated space
Variety of non-laboratory personnel performing testing
Portable devices and kits that are not routinely used in the main laboratory
Documentation by mechanisms different from those used in the main laboratory

## Slide 4: 2013: Common Deficiencies
A number of checklist items are cited frequently during CAP inspections
Most frequently cited deficiencies are similar from one year to the next
This section will review the most commonly cited deficiencies

## Slide 5: Most Commonly Cited Deficiencies in 2013
Chart: 
Chart: 
| Personnel Records | Reagent Storage | Reagent Labeling | Procedure Review | Attestation Page | Procedure Manual | PT Evaluation | Document Control | Activity Menu | Competency
 | 304 | 336 | 364 | 376 | 418 | 427 | 443 | 443 | 543 | 734

## Slide 6: GEN.55500 Competency Assessment
The competency of each person to perform his/her duties is assessed.
Common reasons for citations:
  Must include six competency elements for each non-waived test system (as applicable)
  Must be assessed semi-annually in the first year of duties and annually thereafter

## Slide 7: COM.01200 Activity Menu
The laboratory’s current CAP Activity Menu accurately reflects the testing performed.
Common reasons for citations:
  Some tests missing on activity menu
  Pre-populated activities (waived versus non-waived) are not customized by the laboratory
    Note: Ordering proficiency testing surveys does not automatically add tests to the activity menu

## Slide 8: COM.01700 PT Evaluation
There is ongoing evaluation of PT and alternative assessment results, with prompt corrective action taken for unacceptable results.
Common reasons for citations:
  Missing evaluation of unacceptable results
  Alternative testing not performed twice per year
  Results not reviewed and signed in a timely manner.

## Slide 9: GEN.20375 Document Control
The laboratory has a document control system.
Common reasons for citations:
  Procedures not approved by Laboratory Director
  Multiple versions of procedures available to staff
  Forms and logs not under document control

## Slide 10: COM.10000 Procedure Manual
A complete procedure manual is available at the workbench or in the work area.
Common reasons for citations:
  Procedure manual not in work area
  Procedure manual is incomplete
  Instrument Operator’s Manual in use but not approved for use by Laboratory Director

## Slide 11: COM.01400 Attestation Page
The proficiency testing attestation statement is signed by the laboratory director or designee and the individual performing the testing.
Common reasons for citations:
  Attestation page missing
  Typed names, but no signatures

## Slide 12: COM.10100 Procedure Manual Review
There is documentation of review of all technical policies and procedures by the current laboratory director or designee at least every two years.
Common reasons for citations:
  Review performed by someone not delegated this responsibility
  Review not performed at frequency defined by laboratory (not to exceed two years)

## Slide 13: COM.30300 Reagent Labeling
Reagents, calibrators, controls, and solutions are properly labeled, as applicable and appropriate, with the following elements.
1. Content and quantity, concentration or titer
2. Storage requirements
3. Date prepared or reconstituted
4. Expiration date
Common reasons for citations:
  Modified expiration date per manufacturer not recorded

## Slide 14: COM.30350 Reagent Storage
All reagents and media are stored as recommended by the manufacturer.
Common reasons for citations:
  Room temperature reagents/kits not stored within range specified by manufacturer
  Frozen reagents not stored within range specified by manufacturer
  Frost-free requirements not maintained

## Slide 15: GEN.54400 Personnel Records
Personnel files are maintained on all current technical personnel.
Common reasons for citations:
  Copy of academic diploma or transcript missing
    Note 1: Nursing license is not accepted in lieu of diploma or transcripts
    Note 2: Third party verification of education is not acceptable in lieu of diploma or transcript
  Job description not available

## Slide 16: Top Five Deficiencies – POC Checklist
POC.06900 Competency Assessment
POC.07300 Daily QC
POC.07568 Comparability of Instrument / Method
POC.07037 Documented QC Results - Waived Tests
POC.07512 QC Handling

## Slide 17: Competency Assessment
The 7/29/2013 Checklist Edition broke POC.06900 into 2 different requirements:
  POC.06875 – Competency Assessment for Waived Testing
  POC.06910 – Competency Assessment for Non-waived Testing
Competency Assessment programs for Waived Testing can select which of the 6 elements will be assessed for each test system
Competency Assessment programs for Non-waived testing must include all 6 elements as applicable to each test system
Frequency of competency assessment is the same regardless of testing classification

## Slide 18: POC.06875 Competency Assessment for Waived Testing
There is a documented program to ensure that each person performing waived testing maintains satisfactory levels of competence.
NOTE: Prior to starting patient testing and prior to reporting patient results for new methods or instruments, each individual must have training and be evaluated for proper test performance as required in GEN.55450. After an individual has performed his/her duties for one year, competency must be assessed annually. Retraining and reassessment of employee competency must occur when problems are identified with employee performance. For waived test systems, it is not necessary to assess all 6 elements listed below at each assessment event: the POC program may select which elements to assess. Elements of competency assessment include but are not limited to:
  Direct observations of routine patient test performance,
  Monitoring the recording and reporting of test results,
  Review of test results, QC records, PT results, and maintenance records,
  Direct observation of instrument maintenance and function checks,
  Assessment of previously analyzed specimens, internal blind testing samples or external PT samples; and
  Evaluation of problem-solving skills

## Slide 19: POC.06910 Competency Assessment for Non-waived Testing
There is a documented program to ensure that each person performing non-waived testing maintains satisfactory levels of competence.
  For non-waived test systems, competency using all six elements must be assessed for each individual on each test system during annual and semi-annual assessments, unless an element is not applicable to the test system.
  A TEST SYSTEM is the process that includes pre-analytic, analytic, and post-analytic steps used to produce a test result or set of results.
    A test system may be manual, automated, multi-channel or single use and can include reagents, components, equipment, or instruments required to produce results
    A test system may encompass multiple identical analyzers or devices
    If there are any tests with unique aspects or procedures within the same testing platform (e.g. pretreatment of samples prior to analysis), competency must be assessed as a separate test system

## Slide 20: POC.07300 Daily Quality Control (QC)
Controls are run daily for quantitative and qualitative tests.
  2 levels run daily
  If using electronic or internal controls:
    System must be FDA-cleared or approved and not modified by lab
    System is not classified as high complexity
    Validation study has been performed (after 1/31/2012)
      Must include 20 consecutive days of testing
      If validating multiple identical devices, the 20 consecutive days applies to first device, and the laboratory director will determine the extent of testing for all subsequent devices
    External QC is run for each new lot # or shipment, after major system maintenance or software upgrades and at least every 30 days

## Slide 21: POC.07568 Comparability of Instrument/Method
If the laboratory/POCT program uses more than one instrument/method to test for a given analyte, the instruments/methods are checked against each other at least twice a year for correlation of results.
  Applies to all non-waived testing

## Slide 22: POC.07037 Documented QC Results – Waived Tests
Control results are documented for quantitative and qualitative tests, as applicable.
  Quality control must be performed according to manufacturer instructions
  Testing personnel or supervisory staff must review quality control data on days when controls are run to detect trends/problems
  The laboratory director or designee must review QC data at least monthly
  With respect to internal controls, acceptable control results must be documented, at a minimum, once per day of patient testing for each device.*
  All unacceptable control results must be documented (see below).
  *Acceptable internal control results need not be documented, if (and only if) an unacceptable QC result automatically locks the instrument and prevents release of patient results.

## Slide 23: POC.07512 QC Handling
Control specimens are tested in the same manner and by the same personnel as patient samples.
  Applies to daily QC
  New lot verification can be performed by POC Coordinator or other laboratory staff

## Slide 24: Inspection Process: What Will the Inspector Look For?
Review all centrally located records
Review documentation of maintenance, reagent labeling, testing performed, QC, and results reporting
Ensure that policies and procedures are available to testing personnel
Select several sites to visit based on volume and scope of testing
  High volume: Look for incomplete documentation due to time constraints
  Low volume: Look for incomplete documentation due to lack of familiarity with processes

## Slide 25: Inspection Process cont.
Monitoring of Proficiency Testing (PT)
  Required for all analytes regardless of CLIA complexity classification
  PT must be specific to the analyte and system matrix
  If PT is not available, another mechanism must be defined to assess accuracy and reliability of the system (alternative assessment)
  PT must be performed by, and rotated among, the same individuals who perform the routine testing
  Corrective action must be documented and reviewed for unacceptable results and for non-graded results

## Slide 26: Inspection Process cont.
Quality Control in POCT
  At least two levels of daily QC is required for both waived and non-waived testing
  Internal QC must be validated if used as daily QC
  QC must be verified prior to reporting patient results
  Corrective action must be documented and reviewed
  QC must be handled in the same manner as patient samples and performed by those performing patient testing

## Slide 27: Inspection Process cont.
Maintenance must be performed as required by the manufacturer, documented, and reviewed at least monthly by the laboratory Director or an appropriate designee
Correlations must be performed on non-waived instrumentation at least every 6 months
Calibration verification must be performed on non-waived instrumentation at least every 6 months
Training of testers must be performed initially and then competency must be assessed at appropriate intervals thereafter
Results must be recorded with appropriate reference ranges or reference ranges must be available to practitioners

## Slide 28: What the Inspector Will Look For…
Patient results are not reported if QC is out of range
QC is performed by the same individuals who perform the testing
Corrective action is implemented and recorded
Patient results are reported with reference ranges
Maintenance checks and function checks are performed and documented

## Slide 29: What the Inspector Will Look For…
Proficiency testing is rotated among all testing personnel
Supervisory review of QC, maintenance and patient results
Individual performing the testing is identified on the report
Ongoing training and competency program
All reagents are appropriately labeled
Policies and procedures are available to testing personnel

## Slide 30: You Be the Inspector!
You are inspecting a lab that is doing kit testing on BHCG (urine and serum) and RSV on Bronchial washes and swab specimens types. The lab’s competency assessment groups all kit tests together as one test system and includes 3 out of the 6 elements.
Does this meet the requirements for competency assessment?

## Slide 31: What do you do?
A.) Cite POC.06910 Competency Assessment – Non-waived Testing
  There is a documented program to ensure that each person performing non-waived testing maintains satisfactory levels of competence.
B.) Cite POC.06875 Competency Assessment - Waived Testing
  There is a documented program to ensure that each person performing waived testing maintains satisfactory levels of competence.
  C.) Nothing. The lab is compliant.

## Slide 32: Answer
Cite POC.06910
  The lab has not appropriately identified their test systems. They have grouped non-waived testing with waived testing.
  Non-waived test systems require competency assessment that addresses all six elements of assessment. (They are compliant for their waived testing.)

## Slide 33: Let’s talk about PPT!
Personally performed by provider or midlevel practitioner (such as a physician assistant, nurse practitioner or midwife)
Inspected under the CAP POC checklist:
  If performed under same CLIA number as lab AND
  Lab Director is responsible for competency assessment of provider performing test
NOT inspected by the CAP:
  PPT is performed under the same CLIA number as the laboratory, and
  The institutional medical staff has established the competency of physicians and mid level practitioners through the credentialing process.
  OR PPT is performed under a separate CLIA number
Not the same as PPM

## Slide 34: PPT testing includes
pH, body fluids
Vaginal fluid smears for ferning
Fecal Leukocytes
Gastric biopsy urease
Nasal smears for Eos
Occult blood, fecal and gastric
Pinworm examination
Post-coital mucus examination
KOH prep
Semen analysis, qualitative
Urine dipstick
Urine sediment microscopy
Wet mount prep for the presence of bacteria, fungi, parasites, and human cellular elements.

## Slide 35: Procedures and Policies
Outline PPT testing that may be performed by providers according to their scope of clinical practice
Process for training providers or credentialing process
Assurance that manufacturer instructions are followed
Competency assessment process
Process for documenting results
If providers are credentialed, this may be found in the credentialing documentation

## Slide 36: Let try inspecting PPT!
You are inspecting a POCT site where the physician looks for crystals under the microscope. When you ask for a competency assessment you are told that 6 years ago when the Dr. was hired the lab director deemed him competent.
Does this meet the PPT Requirements?

## Slide 37: What do you do?
A.) Cite POC.09600 PPT Competency Assessment
  There is evidence of competency assessment specific to the type(s) of laboratory testing performed by each provider.
B.) Ask for more information
C.) Nothing. They are compliant since frequency and extent of competency assessment is at the discretion of the laboratory director or the medical staff credentialing process

## Slide 38: Answer
C.) Nothing. They are compliant since frequency and extent of competency assessment is at the discretion of the laboratory director or the medical staff credentialing process

## Slide 39: Resources
Customer Contact Center: 1-800-323-4040
E-mail: accred@cap.org
Tools on the CAP website: e-LAB Solutions > Proficiency Testing and Quality Management > Resources
      PT Troubleshooting Guide
      PT Exception Investigation Checklist
Proficiency Testing Tool Box: www.cap.org: e-lab solutions> Laboratory Accreditation>Proficiency Testing Compliance
Notes: Linda YOU READ
This slide includes contact information and online resources for your laboratory. We hope the sample documents in the Toolkit will be helpful for your laboratory.